- (normative)
Dry heat bioburden reduction proposal - DRD

  - DRD identification

    - Requirement identification and source document

This DRD is called from ECSS-Q-ST-70-57, requirement 5.1c.

    - Purpose and objective

The purpose of the proposal is to describe a bioburden reduction process for a product. It is written by a supplier of a bioburden reduction service and is the response to a bioburden reduction specification.

  - Expected response

    - Scope and content

          - The proposal for dry heat bioburden reduction shall include:

            - Bioburden reduction procedure planned to be used.

            - List and description of equipment needed for applying and controlling the bioburden reduction process.

            - Description of the thermal boundary conditions of the equipment for the selected bioburden reduction temperature.

            - Product specific cycle development.

            - Pre-conditioning for the product.

            - Particular and molecular contamination control before, during and after the bioburden reduction process is applied.

            - Bioburden recontamination control for the product before, during and after the bioburden reduction process is applied.

            - The dry heat bioburden reduction cycle, including heat-up time, bioburden reduction temperature and time as defined by the procedure used (see 5.3.1), cool-down time, and temperature set points for the control of the cycle.

- 1	To achieve this requirement it can be necessary to perform some tests with a thermal representative model equipped with thermocouples under the same conditions as planned for the flight hardware bioburden reduction.

- 2	Two kinds of temperature measurements are typically used in a dry heat bioburden reduction – one kind to monitor the temperature of the product and another kind to monitor and control the process. In case a representative model is used for the cycle development, the temperature measurement on the product can be done during cycle development.

            - Locations to measure the temperature on the product to establish the dry heat bioburden reduction cycle.

- This is usually done based on thermal analysis of the product.

            - The loading pattern of the equipment.

            - Loading pattern specific locations to measure process parameters to control the cycle.

- Typical process parameters to control the cycle are temperature and humidity.

            - Values for process parameters and their tolerances to control the bioburden reduction cycle.

- Proper monitoring and documentation is necessary to allow parametric release of the product, see 5.2.4.

            - The level of purity of the air or other gases used in the process.

            - Environmental conditions and control of the equipment.

- This includes e.g., any level of particulate or molecular contamination control, filtrations systems, use of forced air flow with direction and velocity, use of pumps, pressure level.

            - Deliverables.

            - Quality standards.

    - Special remarks

None.